## Slide 1: New magnetic test bench for measuring closed magnet structures
J. Campmany,
L. Ribó, C. Colldelram, L. Nikitina, F. Becheri, G. Cuní, J. Marcos, V. Massana,
J. V. Gigante, X. Serra, A. Camps, J. Nicolàs, F. Rey, M. Llonch, J. Ladrera,
J. Ferrer, K. Maimouni, D. Calderón
www.cells.es

## Slide 2
The challenge of «closed» structures
Small gap magnets for low emitance rings
H magnets
In-vacuum undulators
Cryoundulators
Any structure that cannot be accessed lateraly

## Slide 3
closed magnetic structures
In-vacuum undulators
Superconducting magnets
Superconducting undulators
Low emmittance rings: MAX IV magnets
(15 mm aperture)
CLIC magnets
(10 mm aperture)

## Slide 4
Small and compact 3D Hall probe
  13.9 mm Wide, 4 mm High, 100 mm long
  Temperature measured with accuracy ±0.01º C
  3 Bell sensors assembled ortogonally
Fixed on a carbon fiber tape. The tape can be attached and deattached and be introduced inside any closed structure with entrance and exit openings
Alignment is achieved using a reference magnet (horizontality) and a cone reference system (positioning)
Whole sistem running into a squared tube acting as anti-chamber or anti-cryostat that can be introduced into a vacuum chamber through the final flanges, allowing Hall probe measurements of cold structures at warm temperature
our proposal

## Slide 5
concept

## Slide 6
Dimensions: 13 x 25 x 2 mm
Weigth: 0.75 g
F.W. Bell Hall sensors, Model GH-700
The probe
Flexible flat cable fixed on stretched tape

## Slide 7
4
Hall sensor attached on a tape tensioned on a C shaped arc structure
The tape is passing through the closed magnetic structure
The arc is moved by a high accuracy motion stage
Key points
Can this arc be moved with an accuracy similar to that of a machining tool? -> stiffness of «C» structure
How much space will we need it? -> measurement range x 3
Is the tape stable after being tensioned? -> vibration studies
L
Hall
sensor
L/
2
L/
2
Arc length 2l
L/
2
L/
2
Length needed for the measurement is 3L
T
T
Y
X
The bench

## Slide 8
Field sentitivity ~10-4 T
Spatial repetitivity ~10-6 m
Measurement range: 3 m longitudinal (for prototype, 1.2 m), 0.25 m horizontal, and 0.1 m vertical
Longitudinal scans allowing vertical and horizontal positioning (3D movement)
Small guidance error on positioning the Hall probe (~ 50·10-6 m)
Very small angular deviations when moving (~50·10-6 rad)
Requirements

## Slide 9
5
Analytical validation
A string under tension has a mode of vibration on the 1st harmonic depending on the vibrating length, the tension and the linear mass
This vibration, if excited, has a clear impact on the accuracy of the Hall probe position
Calculation
L ~ 4 m
Tape cross-section: 24 x 1.4 mm2
Material: pultruded carbon fiber, with density d=1600 kg/m3
Elastic limit: 2800 MPa
T = 7.5 kN -> f = 55 Hz -> stress = 224 MPa
At ALBA, environmental low frequencies are considered < 30 Hz
So, this value is taken as reference when designing

## Slide 10
6
Evaluate fundamental frequency modes of a tensioned tape with the same cross-section as calculated
Experimental set-up: tensioned belt with interferometer
External excitation has been externally induced
The set-up can measure the oscillation amplitude as well as the vibrating frequencies
Empirical validation
Iris
Inerferometer
Carbon
fiber
tape
Result:
Frequencies at 55 and 165 Hz
Amplitude < 5·10-7 m
Test OK: high frequency and low amplitude

## Slide 11
dX, dY, dZ < 0.05 mm
Longitudinal POSITIONING
ERROR
Roll da, Pitch db < 50·10-6 rad
Yaw dg < 100·10-6 rad
Angular POSITIONING ERROR
X, Y, Z = 30·10-6 m
Repeatability
X : ± 0.25 m
Z : ± 0.10 m
Y : ± 1.20 m
Chamber allowance (beam stay clear area) = 0.6 m
Ranges
Y’ = 15 mm/s
Speed
XY
stage
Z
stage
Arc
structure
with
carbon
fiber
tape
Detailed design
PROTOTYPING

## Slide 12
Dimensions
Side View
2600
650
650
4300
Top View
1400
1445 (+125)
600

## Slide 13
X,Y stage
Structure: Al profile
Stretching blocks, one with load cell gauge
Mass: ~ 400 kg
Tape cross-section 16 x 1,4 mm2
Vibrating length: 2.6 m
Density d = 1600 Kg / m3
Pulling force: 0.5 N
Belt dimensioning
Results
Arc structure
Stress = 223 MPa
Security factor = 13
f0 = 71 Hz
Elongation: 0.004 m
Double flexure system on a granite block
Compact design: with a single step Z & pitch
Allowing 100 mm Z range and tilt about 0.2º
Flexures on high Young modulus material
Preloaded guiding system and grinded spindles
Movement for each flexure is encoded
Mass of assembly ~1200 kg
Z stage
X stage has 2 actuators to avoid rotation on vertical axis
Y stage measurement axis: placed on a granite block
Preloaded roller and matched guides, which separation determine vertical accuracy
Mass of the XY stage ~5000 kg

## Slide 14
FEA calculation

## Slide 15
Verification of parts with alignment and metrology group
Big granite block is aligned flat with respect to the floor
Interface plate fitted in top. Slots for guides have been checked
Plate grinding has been corrected several times
Very accurate alignment of linear guide sets
Assembly
XY stage
Z stage
Hollowed granite placed on top of XY stage
Motors have been tested
Arc structure
Stretching system
Carbon fiber is cutted by water
All motors of the assembly have been tested
«C» structure and stretching system have been assembled in parallel
Carbon fiber is stretched up to 10 kN

## Slide 16
Prototype characterization
Straightness of Y displacement
Straightness of Z displacement
Straightness errors
Peak to peak < 20·10-6 m
Guidance errors
Yaw errors < 10·10-6 rad

## Slide 17
Roll rugosity error < 5·10-6 rad
Roll drift error < 50·10-6 rad

## Slide 18
Vibration static
> 50 Hz

## Slide 19
Conclusions
Next steps
Integration of Hall sensor in the bench
Positioning error on Hall probe in X, Y, Z in all range
Repetitivity positioning error on Hall probe in X, Y, Z
Guidance error: straightness, flatness, pitch, yaw and roll errors measured on Hall probe
Roll angle: measurements of repetitivity, hysteresis and evolution
Vibration frequencies for the final assembly
Main granite deformations with Z stage movements
After measurements done, specifications fulfillment is almost guaranteed
Bench can work for very narrow gap closed structures as well as conventional
Main drawback is the long longitudinal space needed as operational range
If the prototype suceeds, a 3 m range bench (9 m total lenght) can be foreseen

## Slide 20
Thanks for your attention!